## Slide 1: Upper Merced Hydroclimate Network
All Gage/Snow Sensor Locks Combo: 264(0)
Created August, 2014-HF
Updated August 10, 2014- HF
Updated March 21, 2017- B.Mocsny

## Slide 2: Merced Lake Ranger Station Cabin Lock requires CG 10 key
Gear Cache updated on 9/2015 N. Ernster
List on Aquarius- site visit date 20150914 or 20150915

## Slide 3: Merced River ABV HSC
Stream Gage Enclosure:
-CR1000 data logger(PB address=471)
-CS547A Conductivity Interface
-RF400 Local Radio(set to .4 mAmp-8 sec
transmission, 8/14/2014)
-CSI Yagi directional local antennae
-data cable connection to enclosure via 1 ½ inch
coupler
-enclosure door falls off when open, align hinge
grommets to reattach and close door
Power System:
-20 Watt Solarex solar panel
-PS12 Solar Controller w/7AH Battery
-External 55 AH Battery in lower ammo can.
-Power cable from solar panel to solar
controller = 40 ft.
-power cable connection to enclosure = ½ inch
conduit coupler.

## Slide 4: Merced River ABV HSC
Sensors:
“Alpine Stilling Tube”(CS450 PT & CS547 Cond)
Installed Aug.7,2014
CS451 Pressure Transducer/Stage Recorder
-Cable Length: 30 ft.
-Serial #: 20010250
CS547 Conductivity Sensor
-Cable Length: 30 ft.
-Calibration K: 1.380
-Serial #:6425
Offset applied to CS450 to Match Gage Height on 9/22/2014: +0.03 ft.

## Slide 5: Merced River ABV HSC
Sensors:
“Original Hydroclimate Gage” Installed July 2001
Druck PT (still running)
-Cable Length: 100 ft.
-Serial Number: NA
CS 547 Cond (in situ, out of service)
-Cable cut at lower enclosure on 8/7/2014
Offset applied to Druck to Match Gage Height on 9/22/2014: -0.02 ft.

## Slide 6: Merced River ABV HSC
Datum:
Tape Down Bolt
(Installed August 7, 2014)
Measure from center of “X” inscribed
on face of bolt.
Gage Height = 10ft. – / + Tape Down
If inaccessible due to high flow, use
“Original Post Tape Down”
* Measure Gage Height from both Tape Down Bolt and “Original Post Tape Down” *
Alpine
Stilling
Tube

## Slide 7: Merced River ABV HSC
Datum:
“Original Post Tape Down”
Measure from top edge of post.
Gage Height = 12.1 ft. – / + Tape Down
10 ft. on Tape Down Bolt = 7.90 on “Original Post Tape Down”
* Measure Gage Height from both Tape Down Bolt and “Original Post Tape Down” *
You should get the same Gage Height from both Tape Down locations!

## Slide 8: Merced River ABV HSC
Datum:
“Original Post Tape Down”
Tape Down Bolt
Sensors:
“Alpine Stilling Tube”

## Slide 9: Merced River ABV HSC
Datum:
“Original Post Tape Down”
Tape Down Bolt
Sensors:
“Alpine Stilling Tube”
Section Control
Tape Down Bolt
Point of Zero Flow(PZO) / Gage Height(GHZO) = Gage Height - Depth of Water over Section Control
Gage Height of Zero Flow = (insert GH) – (insert Water Depth) = (insert PZO) on (insert date)

## Slide 10: Merced River ABV HSC
Meteorology Sensors:
Hobo Temp/RH
Tidbit Temp
Barologger Enclosure/Sensor
To Stream Gage, 30 ft. SW

## Slide 11: Merced River ABV HSC
Snow/Met Sensors:
Judd Snow Depth
-cable length= 15 ft.
-sensor height: 248 inches(8/14/14)
-serial #: NA
CS215-L Temperature & Relative Humidity
-cable length= 30 ft.
-serial #: E10466
-sensor height: 15 ft.
Out of service sensors to be removed September 2014:
-sonic anemometer
-solar radiation
-Viasala Temp/RH
To Stream Gage, 300 ft. SE

## Slide 12: Merced River ABV HSC
Snow Sensor Enclosure;
-CR1000 Datalogger
-RF400 Radio
-CSI TX312 HDR GOES Transmitter
-CSI Omnidirectional White-Wand Local Antennae
-GOES YAGI Directional Antennae(Azimuth=191°)
Power System:
40 Watt BP solar panel
ProStar-15 solar controller
CS P100 controller(to power GOES Transmitter)
80 AH Battery

## Slide 13: Merced River ABV HSC
CSI TX312 HDR GOES Transmitter
Sat Configuration:
GOES ID: c2b033ca
Timed Channel: 62
Timed Bit Rate: 300
Timed Interval: 00:01:00:00
First Timed Transmission: 00:12:30
Timed Tx Window: 10
Timed Data Format: A
Random Channel: 0
Random Bit Rate: 100
Random Randomizing Interval: 20
Random Randomizing Percentage: 50
Random Repeat Count: 5
Random Data Format: A
Random Msg Counter: N
GPS Position Fix Interval: 00:00:00

## Slide 14: Merced River BLW Lyell Fk.
Datum:
Tape down to water surface from bridge v-notch
Solinst:
-Solinst Black Sealed PT inside pancake PT mount.
(fixed inside w/zip ties 8/14/2014)

## Slide 15: Merced River BLW Lyell Fk.
Locating the Meteorology Sensor:
HOBO Temp/RH
HOBO Tidbit Temp
Walk ~150m downstream from bridge
On RL.
Look into the forest for a large tree and
Rootwad down. Walk this direction.

## Slide 16: Merced River BLW Lyell Fk.
Locating the Meteorology Sensor:
HOBO Temp/RH
HOBO Tidbit Temp
This is the fallen tree with giant
rootwad.
This picture is taken standing at the tree with the sensors on it looking towards the rootwad.

## Slide 17: Merced River BLW Lyell Fk.
Locating the Meteorology Sensor:
HOBO Temp/RH
HOBO Tidbit Temp
Towards Trail
Towards Rootwad

## Slide 18: Merced Peak Fk.
Stream Gage Enclosure:
-CR800 data logger (PB address=222)
-CS547A Conductivity Interface
-data cable connection to enclosure via 1 ½ inch
coupler reduced to ¾ inch conduit coupler, ¾ inch
conduit
Power System:
-20 Watt Solarex solar panel
-PS1000 Solar Controller w/7AH Battery
-Power cable from solar panel to solar
controller = 50 ft.
-power cable connection to enclosure = ½ inch
conduit coupler, ½ inch conduit.

## Slide 19: Merced Peak Fk.
Sensors:
“Piton PT Mount !”(CS450 PT)
Installed Aug.6, 2014
CS450-L Pressure Transducer/Stage Recorder
-Cable Length: 50 ft.
-Serial #: 70010415 Cal Date: 20101022
CS547 Conductivity Sensor(to be installed Sept. 2014
-Cable Length: 50 ft.
-Calibration K: 1.410
-Serial #:6659
Offset to Match Gage Height on 8/6/2014:
+7.85 ft.

## Slide 20: Merced Peak Fk.
Sensors:
“Simple Solinst Stake Mount”
Installed Aug.6, 2014
Solinst Black Sealed Pressure Transducer
Stainless steel set screw fixes through top eyelet on Solinst

## Slide 21: Merced Peak Fk.
Datum:
Tape Down Bolt
(Installed August 6, 2014)
Measure from center of “X” inscribed
on face of bolt.
Gage Height = 10ft. – / + Tape Down

## Slide 22: Merced Peak Fk.
Datum:
Tape Down Bolt
(Installed August 6, 2014)
Measure from center of “X” inscribed
on face of bolt.
Gage Height = 10ft. – / + Tape Down
“Piton PT Mount !”
“Simple Solinst Stake Mount”
Tape Down Bolt
Control

## Slide 23: Merced Peak Fk.
Control
(Pre-“Piton PT Mount !” installation)
Point of Zero Flow(PZO) / Gage Height(GHZO) = Gage Height - Depth of Water over Control
Gage Height of Zero Flow = 9.19 - .61 = 8.58 on August 6, 2014

## Slide 24: Merced Peak Fk.
Meteorology Sensors:
Hobo Temp/RH
Tidbit Temp

## Slide 25: Lyell Fk. Merced River ABV High Trail
Stream Gage Enclosure:
-CR1000 data logger(PB address=296)
-CS547A Conductivity Interface
-data cable connection to enclosure via 1 ½ inch
coupler reduced to ¾ inch conduit coupler, ¾ inch
conduit
Power System:
-15 Watt Solarex solar panel
-PS100 Solar Controller w/7AH Battery
-Power cable from solar panel to solar
controller = 4 ft.
-power cable connection to enclosure = ½ inch
conduit coupler.

## Slide 26: Lyell Fk. Merced River ABV High Trail
Sensors:
“Alpine Stilling Tube”(CS451-L PT & CS547 Cond)
Installed Aug.5, 2014
CS451-L Pressure Transducer/Stage Recorder
-Cable Length: 50 ft.
-Serial #: 20010336 Cal Date: 20131127
CS547 Conductivity Sensor
-Cable Length: 50 ft.
-Calibration K: 1.383
-Serial #: __61 (difficult to read)
Old Stage Post TD + 0.27ft=New Stage TD Bolt.
Offset to Match Gage Height on 8/5/2014:
+7.13 ft.

## Slide 27: Lyell Fk. Merced River ABV High Trail
Sensors:
“Simple Solinst Stake Mount”
Installed: Unknown date
Solinst Black Sealed Pressure Transducer
Stainless steel set screw fixes through top eyelet on Solinst
“Simple Solinst Stake Mount”

## Slide 28: Lyell Fk. Merced River ABV High Trail
Datum:
Tape Down Bolt
(Installed August 5, 2014)
Measure from center of “X” inscribed
on face of bolt.
Gage Height = 10ft. – / + Tape Down

## Slide 29: Lyell Fk. Merced River ABV High Trail
Control
Point of Zero Flow(PZO) / Gage Height(GHZO) = Gage Height - Depth of Water over (deepest part of) Control
Gage Height of Zero Flow = 7.74 - .9 = 6.84 on August 5, 2014
Control

## Slide 30: Lyell Fk. Merced River ABV High Trail
Meteorology Sensors:
Hobo Temp/RH
Tidbit Temp

## Slide 31: Remember the Side Burns: The integrity of each data point collected fortifies or degrades a legacy that lives on…